BRAUN

Shelly Power Logger

Hardware Description Document  •  v1.1

Portable 230 V power logger

Shelly Plug S MTR Gen3  +  ESP32-WROOM-32  +  DS3231 RTC  +  MicroSD

Author: Jan Pfrang

jan.pfrang@delonghigroup.com

Date: May 28 2026

# 1.  System Overview

The BRAUN Shelly Power Logger is a portable, self-contained instrument for measuring and recording the power consumption of 230 V / 16 A appliances. It requires no external Wi-Fi router or internet connectivity — the entire system forms its own local network.

## 1.1  Architecture

The system consists of three physical units:

- Shelly Plug S MTR Gen3 — CE-marked smart plug measuring voltage, current, and active power at ~1 Hz. Runs a custom script (shelly_push.js) that pushes JSON measurements over Wi-Fi to the ESP32 every second.

- ESP32-WROOM-32 — acts as the Wi-Fi access point (softAP), web server, RAM buffer, SD card writer, and real-time clock host. Serves a browser-based dashboard over HTTP.

- DS3231 ZS042 RTC module — provides accurate wall-clock timestamps for log entries, connected via I²C.

## 1.2  Data flow

Shelly script (shelly_push.js)  →  POST /api/shelly_push  →  ShellyClient::ingest()  →  Logger::pollIfDue()  →  RAM ring buffer  →  SD card (CSV)

## 1.3  Key specifications

| Parameter | Value |
| --- | --- |
| Mains supply | 230 V AC, 50 Hz |
| Max load current | 10 A (2500 W) |
| Measurement cadence | 1 s (default); adjustable 1 / 2 / 5 / 10 / 30 s |
| SD flush interval | Every 10 s |
| RAM buffer depth | 64 samples (~64 s reserve) |
| Log format | CSV: time_ms, voltage_V, power_W, pf_apparent |
| Wi-Fi | 802.11 b/g/n softAP — SSID: PZEM_Logger |
| Web interface | http://192.168.4.1 or http://braun_PZEM.local |
| RTC accuracy | DS3231 ±2 ppm (±63 s/year) |
| LED: 1 Hz blink | System healthy |
| LED: 5 Hz blink | Error (Shelly silent > 3 s, or SD fault) |

# 2.  Component List

| Component | Description | Notes |
| --- | --- | --- |
| Shelly Plug S MTR Gen3 | Smart plug, 230 V / 10 A, built-in energy meter | CE-marked; shelly_push.js installed |
| ESP32-WROOM-32 | Dual-core 240 MHz, 4 MB flash, Wi-Fi + BT | Arduino framework |
| DS3231 ZS042 module | I²C RTC, TCXO-compensated, CR2032 backup | Address 0x68; see battery note |
| MicroSD card | FAT32-formatted, any capacity | Writes CSV log.csv |
| 230 V → 5 V PSU | Mains-rated isolated supply, ≥ 500 mA | Must be safety-certified |
| Status LED | Single LED + current-limiting resistor | Connected to GPIO 32 |
| Reset button | Tactile pushbutton | GPIO 33, INPUT_PULLUP (future use) |
| Balancing resistors | 2 × 100 kΩ | For 2S supercap stack |
| Voltage divider for P35 – 9V to ADC ESP32 | 1 × 180 kΩ, 1 × 47 kΩ |  |

2.1  Complete Circuit Schematic

The schematic below shows the complete wiring of the logger. The LED circuit is active-HIGH configuration: GPIO 32 sources current through R3 (220 Ω) and the LED to GND, matching the firmware logic in StatusLed.h.

WRONG/not complete IN SCHEMATIC:

C1 = C2 = 1F

NTC is 10Ω while cold

Figure 1  —  Complete circuit schematic Rev 1.1. All pin labels inside component boxes. LED wiring corrected to active-HIGH (P32 → R3 220Ω → LED → GND). R1/R2 voltage divider supplies ADC on P35.

# 3.  ESP32 Pin Assignment

The table and diagram below document every GPIO used in this design. GPIOs 16 and 17 were used for the PZEM UART in the previous hardware revision; they are now free after the switch to the wireless Shelly sensor.

## 3.1  Pin assignment diagram

Figure 2  —  ESP32-WROOM-32 complete GPIO assignment. Faded entries are unconnected.

## 3.2  Pin assignment table

| GPIO / Pin | Connected to | Category |
| --- | --- | --- |
| 3.3V | RTC VCC + board supply rail | Power |
| GND | RTC GND + board ground | Power |
| 5V / VIN | 5 V input from PSU | Power |
| GPIO 14 | SD card MOSI (SPI VSPI) | SD SPI |
| GPIO 25 | SD card CS | SD SPI |
| GPIO 26 | SD card MISO (SPI VSPI) | SD SPI |
| GPIO 27 | SD card CLK (SPI VSPI) | SD SPI |
| GPIO 21 | RTC DS3231 SDA (I²C) | RTC I2C |
| GPIO 22 | RTC DS3231 SCL (I²C) | RTC I2C |
| GPIO 32 | Status LED (output) | UI |
| GPIO 33 | Reset button (INPUT_PULLUP) | UI |
| GPIO 35 | Supply voltage ADC (input only) | ADC |
| GPIO 16 | Free | Free |
| GPIO 17 | Free | Free |
| GPIO 1 | UART0 TX (Serial / debug) | System |
| GPIO 3 | UART0 RX (Serial / debug) | System |
| GPIO 0 | Boot strapping — do not drive | System |
| GPIO 2 | Boot strapping — do not drive | System |
| GPIO 15 | Free | Free |
| GPIO 4 | Free | Free |
| GPIO 5 | Free | Free |
| GPIO 18 | Free | Free |
| GPIO 19 | Free | Free |
| GPIO 23 | Free | Free |
| GPIO 34 | Free (input only, no pull) | Free |
| GPIO 36 | Free (input only, no pull) | Free |
| GPIO 39 | Free (input only, no pull) | Free |

# 4.  DS3231 RTC Wiring

The DS3231 ZS042 module connects to the ESP32 via the hardware I²C bus (Wire bus 0). GPIO 21 and GPIO 22 are the dedicated SDA and SCL pins on the ESP32-WROOM-32 and include internal pull-up support.

Figure 3  —  DS3231 ZS042 to ESP32 wiring. All signals are 3.3 V logic.

## 4.1  Connection table

| DS3231 pin | ESP32 pin | Signal | Notes |
| --- | --- | --- | --- |
| SDA | GPIO 21 | I²C data | 4.7 kΩ pull-up recommended |
| SCL | GPIO 22 | I²C clock | 4.7 kΩ pull-up recommended |
| VCC | 3.3V | Power | Do NOT connect to 5V |
| GND | GND | Ground |  |
| 32K | — | 32 kHz output | Leave unconnected if unused |
| SQW | — | Square wave / INT | Leave unconnected if unused |

## 4.2  Battery note (ZS042 module)

Many ZS042 boards include a trickle-charging circuit intended for LIR2032 rechargeable cells. If a non-rechargeable CR2032 is installed, the charging resistor (R1, typically 200 Ω) must be removed or the VCC→BAT trace cut to prevent overcharging.

# 5.  Supercapacitor Balancing (Optional)

If two supercapacitors are stacked in series across the 9 V supply rail, passive balancing resistors are required to equalise the voltage across each cell and prevent either from exceeding its rated voltage.

Figure 4  —  Series supercap stack with passive balancing resistors R1 and R2.

## 5.1  How it works

R1 (100 kΩ) connects from the positive supply rail (+9 V) down to the midpoint — the junction between C1 and C2. R2 (100 kΩ) connects from the midpoint down to GND. Together, R1 and R2 form a voltage divider that forces the midpoint to +4.5 V, so each capacitor sees at most half the supply voltage regardless of individual leakage mismatch.

## 5.2  Design values

| Parameter | Value | Formula | Notes |
| --- | --- | --- | --- |
| Supply voltage | 9 V | — | Across C1 + C2 in series |
| Voltage per cap | 4.5 V | 9 V / 2 | Must be ≤ cap voltage rating |
| R1 = R2 | 100 kΩ | Equal divider | Higher R = lower bleed current |
| Bleed current | ≈90 μA | 9V / 100kΩ | Negligible power loss |
| Power in R1/R2 | ≈0.4 mW each | V²/R | 4.5²/100k |

Important: Verify the individual capacitor voltage rating before powering up. Common coin-type supercaps are rated 2.7 V — a 2S stack at 9 V would overvolt them. Use capacitors rated ≥ 5 V per cell, or reduce the supply voltage accordingly.

# 6.  Firmware Module Overview

| File | Responsibility |
| --- | --- |
| Config.h | All pin definitions, timing constants, Wi-Fi credentials, SD path |
| ShellyClient.h | Receives and validates JSON push from Shelly; watchdog timer |
| Logger.h | RAM ring buffer, SD card writes, poll scheduling, power threshold |
| WebPortal.h/.cpp | HTTP server, all web pages and API endpoints, captive portal, mDNS |
| StatusLed.h | Non-blocking LED blink — 1 Hz OK, 5 Hz error |
| shelly_push.js | Runs ON the Shelly device; reads meter and POSTs JSON every 1 s |
| Shelly_ESP32_Logger.ino | Main sketch: object instantiation, setup(), loop() |

## 6.1  Required Arduino libraries

- ArduinoJson  by Benoit Blanchon  (≥ v6.x)  —  JSON parse/serialise

- RTClib  by Adafruit  —  DS3231 driver

- SD  (built-in ESP32 core)

- WiFi, WebServer, ESPmDNS, DNSServer  (built-in ESP32 core)

End of document